## How to Use
How to Use:
This spreadsheet tracks inventory and calibration dates. Rows are color coded by calibration status.
Some Column Notes:
A product is distinct from an item in that there may be multiple items of a given product. An example is a survey meter. We may have two survey meters of the same type, but each survey meter (item) has its own serial number.
- On license: Whether or not the item is under license.
- Vendor doc: Whether or not we have vendor documentation on our network drive.
- TG-40 req for cal: Calibration requirement according to AAPM Task Group Report 40. Must be initial use (case insensitive) or positive integer number of years.
- Date last cal: Date of last calibration of the item. Must be not req'd (case insensitive) or past date (including today).
- Cal due date: Due date of the item's next calibration. Calculation based on TG-40 req for cal and Date last cal.
- Preparing for cal: Whether or not we are currently preparing the item for sending out for calibration. If no due date, must be blank.
- Out for cal: Whether or not the item is currently out for calibration. If no due date, must be blank. If preparing for calibration, must be FALSE.
Technical/Implementation Details:
The column Neither prepping, out, nor late is not for user use so should remain hidden. It is only used for conditional highlighting (coloring) rules.
The sheet is protected to restrict editing to data entry only. There is no password, though, so simply temporarily unprotect the sheet to make any necessary modifications.

## Inventory
Color Legend | [13 empty] | HIDE THIS COLUMN
Preparing for calibration
Currently out for calibration
Late for calibration
Next due
Type | Manufacturer | Product | Description | S/N | Location | On license | Vendor doc | TG-40 req for cal | Date last cal | Cal due date | Preparing for cal | Out for cal | Notes | Neither prepping, out, nor late
Ion chamber | PTW | TN23343 | Markus plane-parallel ion chamber | 2905 | Elekta vault | FALSE | TRUE | 2 | 2020-08-26 | 2022-08-26 12:00 | FALSE | FALSE |  | 2022-08-26 12:00
Ion chamber | PTW | N23343 | Intracavity ion chamber | 3704 | Hot lab | FALSE | TRUE |  | 1999-04-08 |  |  |  | rarely used
Ion chamber | PTW | N30001 | Farmer cylindrical ion chamber | 619 | Elekta vault | FALSE | TRUE | 2 | 2020-08-25 | 2022-08-25 12:00 | FALSE | FALSE |  | 2022-08-25 12:00
Ion chamber | Standard Imaging | A1SL | Exradin Slimline Miniature Shonka Chamber | XW0531111 | Tomo vault | FALSE | TRUE | 2 | 2021-04-28 | 2023-04-28 12:00 | FALSE | FALSE | "light" | 2023-04-28 12:00
|  |  |  | XW052833 | Tomo vault | FALSE | TRUE | 2 | 2020-03-16 | 2022-03-16 12:00 | FALSE | FALSE | "dark"
Ion chamber | Standard Imaging | A17 | Exradin Slice Therapy ion chamber | XAE080647 | Tomo vault | FALSE | TRUE
Ion chamber | Fluke | 30-404 | Holt Parallel Plate Ion Chamber | MPPK 450 3/1 | Block room | FALSE | FALSE | 2 | Not req'd |  |  |  | out of service
Electrometer | Standard Imaging | MAX-4000 Plus |  | DEMO2 | E1 vault | FALSE | TRUE | 2 | 2019-05-16 | 2021-05-15 12:00 | FALSE | TRUE
Electrometer | Standard Imaging | Tomoelectrometer |  | L060401 | Block room | FALSE | TRUE | 2 | 2020-01-20 | 2022-01-19 12:00 | FALSE | FALSE
Electrometer | Fluke | 37-705 | VeriDose Electrometer for Diodes (Neg Bias) | 91697 | Elekta vault | FALSE | TRUE
Survey meter | Victoreen | 450P | Handheld ion chamber survey meter | 3829 | Tomo console | FALSE | TRUE | 1 | 2021-03-28 | 2022-03-28 06:00 |  | FALSE
Survey meter | Victoreen | 451P | Handheld ion chamber survey meter | 3949 | Physics cabinet | TRUE | TRUE | 1 | 2021-09-07 | 2022-09-07 06:00 |  | FALSE | CE 6839 - date good | 2022-09-07 06:00
Survey meter | Ludlum | 2241-3 | Handheld chamber survey meter electrometer | 248122 | Physics office | FALSE | TRUE | 1 | 2021-02-01 | 2022-02-01 06:00 |  | FALSE | electr. Cal. included in the probe cal
Survey meter | Ludlum | 44-9 | Pancake G-M Detector (alpha, beta, gamma,x) | PR266168 | Physics office | FALSE | TRUE | 1 | 2021-02-01 | 2022-02-01 06:00 |  | FALSE
Survey meter | Ludlum | 44-2 | Cylindrical NaI Scintillation Detector (gamma,x) | PR266834 | Physics office | FALSE | TRUE | 1 | 2021-02-01 | 2022-02-01 06:00 |  | FALSE
Survey meter | Ludlum | 133-7 | Cylindrical G-M Detector (gamma,x) | PR247807 | Physics office | FALSE | TRUE | 1 | 2021-02-01 | 2022-02-01 06:00 |  | FALSE
Scanning detector | Scanditronix | DEB050 | Stereotactic Field Detector diode | 1956 | Physics office | FALSE | TRUE
Scanning detector | Scanditronix | DEB022 | Reference Field Detector diode | 3823 | Physics office | FALSE | TRUE
Scanning detector | Scanditronix | DEB012 |  | 3912 | Physics office | FALSE | TRUE
Scanning detector | Scanditronix | DEB002 | Electron Field Detector diode (Blue) | 3988 | Physics office | FALSE | TRUE
Scanning detector | PTW | TN60019 | Diamond Detector |  | Hot lab | FALSE | TRUE | 2 | 2019-01-31 | Not req'd
In-vivo dosimeter | Fluke |  | VeriDose Diode 5-11 MV (Neg Bias) | 95788 | Elekta vault | FALSE | TRUE
In-vivo dosimeter | Fluke |  | Veridose Diode 18-25 MV (Neg Bias) | 1342006 | Elekta vault | FALSE | TRUE
In-vivo dosimeter | Fluke |  | Veridose Diode - 5 - 25 MeV Electron (Neg Bias) | 106634 | Elekta vault | FALSE | TRUE
In-vivo dosimeter | PTW |  | Rectal Intracavitary Ion Chamber | 3704 | Block room | FALSE | FALSE | 2 |  |  |  |  | out of service
Room monitor | Fluke | 05-434 Primalarm | Teletherapy rad monitor | 108070 | Tomo console | FALSE | FALSE
Room monitor | Fluke | 05-433 Primalert 10 | Remote alarm | 108079 | Tomo vault | FALSE | TRUE
Room monitor | Fluke | 05-434 Primalarm | Teletherapy rad monitor | 108069 | Elekta console | FALSE | FALSE
Room monitor | Fluke | 05-433 Primalert 10 | Remote alarm | 108070 | Elekta vault | FALSE | TRUE
Room monitor |  |  | Radiation Monitor | 340705 | E1 vault | FALSE | FALSE | [5 empty] | *s/n may be incorrect
Room monitor |  |  | Radiation Monitor |  | E2 vault | FALSE | FALSE
Room monitor |  |  | Radiation Monitor |  | Tomo vault | FALSE | FALSE
Thermometer |  |  | Type K Thermocouple Thermometer | 440883 | Tomo vault | FALSE | TRUE | Initial use
Thermometer |  |  | Mercury Thermometer |  | Tomo vault | FALSE | FALSE | Initial use
Barometer |  |  | Aneroid Barometer |  | Tomo console | FALSE | TRUE | 4
Barometer |  |  | Altimeter/Barometer |  | Tomo console | FALSE | TRUE | [5 empty] | out of service
Barometer |  |  | Mercury Barometer |  | Physics office | FALSE | TRUE | Initial use
Film equipment | Vidar | DosimetryPRO Advantage | Film scanner | 330025 | Physics office | FALSE | TRUE
Daily/monthly QA equipment | Sun Nuclear | Daily QA 3 | Linac Daily QA Checker | 5283306 | Elekta vault | FALSE | TRUE
Daily/monthly QA equipment | Sun Nuclear | Tomodose | Tomo daily QA | 5498918 | Tomo vault | FALSE | TRUE
Daily/monthly QA equipment | Sun Nuclear | IC PROFILER | Profiler Detector Array |  | Block room | FALSE | TRUE
Daily/monthly QA equipment |  |  | Profiler Power Supply |  | Block room | FALSE | TRUE
Daily/monthly QA equipment |  | PDI 3.0 port | IC Profiler uses PDI 3.0 port |  | Hot lab | FALSE | TRUE
Annual QA equipment |  |  | BP1 3D Water Tank |  | Hot lab | FALSE | FALSE
Annual QA equipment |  |  | Base w/ wheels for tank |  | Hot lab | FALSE | FALSE
Annual QA equipment |  |  | CCU for myQAAccept current version |  | Hot lab | FALSE | FALSE
Annual QA equipment |  |  | CCU for old OmniPro IBA software |  | Physics mid room | FALSE | FALSE
Annual QA equipment |  |  | 2D water tank fits A1SL chamber |  | Tomo closet | FALSE | FALSE
Annual QA equipment |  |  | Control unit for water tank |  | Tomo closet | FALSE | FALSE
Annual QA equipment | Standard Imaging | TomoScanner v2.0 | CD w/ software |  | Physics office | FALSE | FALSE
Linac | Elekta | Synergy | E1 | 151769 | E1 vault |  | TRUE
Linac | Elekta | Infinity | E2 | 154169 | E2 vault |  | TRUE
Software | Varian | MOSAIQ 4.0 | MOSAIQ EMR | #N/A |  |  | TRUE
Linac | Accuray | TomoTherapy Hi-Art | Tomo | 110227 | Tomo vault |  | TRUE
Software | MIM | MIM 6.9.3 | MIM
Formulas:
Cal due date = IF(AND(ISNUMBER(Inventory[TG-40 req for cal]),ISNUMBER(Inventory[Date last cal])),Inventory[Date last cal]+Inventory[TG-40 req for cal]*365.25,"")