## Data
Background Info:
Recently, IMRT QA started consistently failing on E2 even if the same plan passed very well on E1. We had a few different hypotheses: 1. E2 needed recalibrating in Delta4. We had been using a single Radiation Device, ELEKTA, for both E1 and E2. ELEKTA is the main Elekta machine used in RayStation, and it comes from E1 measurements. 2. A recent E2 energy adjustment had unintended consequences. We created a new Radiation Device E2 in Delta4 and re-ran several QA plans on different machines and using different Radiation Devices. This spreadsheet compares GPRs across body sites, time, machine, and Radiation Device, and beam vs. plan. All data uses a dose threshold of 10 percent, a DTA pass criterion of 3 mm, a dose deviation pass criterion of 2 mm, a GPR action level of 90 percent, and a GPR tolerance level of 95 percent.
Date | Pt | Case | Plan | Body Site | Tx Technique | Machine | TreatmentMachineName / Radiation Device | Beam 1 GPR | Beam 2 GPR | Plan GPR | Before/After Energy Adjustment
2022-05-09 | 1 | Thorax | R Lung | Lung | VMAT | E1 | ELEKTA | 97.2 | 98.3 | 95.8 | After
2022-05-09 | 1 | Thorax | R Lung | Lung | VMAT | E2 | ELEKTA | 88.7 | 96.5 | 87 | After
| 1 | Thorax | R Lung | Lung | VMAT | E1 | E1
| 1 | Thorax | R Lung | Lung | VMAT | E2 | E2
2022-05-09 | 2 | SBRT L Lung | SBRT L Lung | Lung | SBRT | E1 | ELEKTA | 100 | 100 | 100 | After
2022-05-09 | 2 | SBRT L Lung | SBRT L Lung | Lung | SBRT | E2 | ELEKTA | 100 | 100 | 100 | After
| 2 | SBRT L Lung | SBRT L Lung | Lung | SBRT | E1 | E1
| 2 | SBRT L Lung | SBRT L Lung | Lung | SBRT | E2 | E2
2022-05-09 | 3 | Thorax | SBRT LML | Lung | SBRT | E1 | ELEKTA | 99.5 | 100 | 99.8 | After
2022-05-09 | 3 | Thorax | SBRT LML | Lung | SBRT | E2 | ELEKTA | 100 | 100 | 100 | After
| 3 | Thorax | SBRT LML | Lung | SBRT | E1 | E1
| 3 | Thorax | SBRT LML | Lung | SBRT | E2 | E2
2022-05-06 | 4 | Pelvis | Rectum | Rectum | VMAT | E1 | ELEKTA | 99.3 | 99.8 | 99.9 | After
2022-05-06 | 4 | Pelvis | Rectum | Rectum | VMAT | E2 | ELEKTA | 95.2 | 92.6 | 82.9 | After
| 4 | Pelvis | Rectum | Rectum | VMAT | E1 | E1
| 4 | Pelvis | Rectum | Rectum | VMAT | E2 | E2
2022-05-05 | 5 | VMAT Rectum | VMAT Rectum | Rectum | VMAT | E1 | ELEKTA | 99.3 | 99.8 | 98.7 | After
| 5 | VMAT Rectum | VMAT Rectum | Rectum | VMAT | E2 | ELEKTA
| 5 | VMAT Rectum | VMAT Rectum | Rectum | VMAT | E1 | E1
| 5 | VMAT Rectum | VMAT Rectum | Rectum | VMAT | E2 | E2
2022-05-05 | 5 | VMAT Rectum | VMAT Rectal Bst | Rectum | VMAT | E1 | ELEKTA | 98.5 | 99.2 | 98.5 | After
| 5 | VMAT Rectum | VMAT Rectal Bst | Rectum | VMAT | E2 | ELEKTA
| 5 | VMAT Rectum | VMAT Rectal Bst | Rectum | VMAT | E1 | E1
| 5 | VMAT Rectum | VMAT Rectal Bst | Rectum | VMAT | E2 | E2
2022-04-15 | 6 | Thorax | R Lung Retx | Lung | VMAT | E1 | ELEKTA | 99.7 | 99.7 | 99.7 | Before
| 6 | Thorax | R Lung Retx | Lung | VMAT | E2 | ELEKTA
| 6 | Thorax | R Lung Retx | Lung | VMAT | E1 | E1
| 6 | Thorax | R Lung Retx | Lung | VMAT | E2 | E2
2022-05-03 | 7 | Pelvis | Rectum | Rectum | VMAT | E1 | ELEKTA | 99 | 99.9 | 99 | After
| 7 | Pelvis | Rectum | Rectum | VMAT | E2 | ELEKTA
| 7 | Pelvis | Rectum | Rectum | VMAT | E1 | E1
2022-05-10 | 7 | Pelvis | Rectum | Rectum | VMAT | E2 | E2 | 97.6 | 96.7 | 93.6 | After
2022-05-03 | 7 | Pelvis | Rectal Boost | Rectum | VMAT | E1 | ELEKTA | 96.4 | 99.7 | 98.9 | After
| 7 | Pelvis | Rectal Boost | Rectum | VMAT | E2 | ELEKTA
| 7 | Pelvis | Rectal Boost | Rectum | VMAT | E1 | E1
2022-05-10 | 7 | Pelvis | Rectal Boost | Rectum | VMAT | E2 | E2 | 98.9 | 97.4 | 95.7 | After
2022-05-03 | 8 | Abdomen | VMAT Epigastric | Abdomen | VMAT | E1 | ELEKTA | 99.1 | 98.9 | 94.3 | After
| 8 | Abdomen | VMAT Epigastric | Abdomen | VMAT | E2 | ELEKTA
| 8 | Abdomen | VMAT Epigastric | Abdomen | VMAT | E1 | E1
| 8 | Abdomen | VMAT Epigastric | Abdomen | VMAT | E2 | E2
| 9 | HNK | VMAT R Lung | Lung | VMAT | E1 | ELEKTA
2022-04-26 | 9 | HNK | VMAT R Lung | Lung | VMAT | E2 | ELEKTA | 99.8 | 99.1 | 99.8 | Before
2022-05-10 | 9 | HNK | VMAT R Lung | Lung | VMAT | E2 | ELEKTA | 97.5 | 96.3 | 97.3 | After
| 9 | HNK | VMAT R Lung | Lung | VMAT | E1 | E1
| 9 | HNK | VMAT R Lung | Lung | VMAT | E2 | E2
2022-04-13 | 10 | VMAT Rectum | VMAT Rectum | Rectum | VMAT | E1 | ELEKTA | 99.2 | 99.5 | 98.9 | Before
2022-05-10 | 10 | VMAT Rectum | VMAT Rectum | Rectum | VMAT | E2 | ELEKTA | 99.7 | 95.6 | 95.6 | After
| 10 | VMAT Rectum | VMAT Rectum | Rectum | VMAT | E1 | E1
| 10 | VMAT Rectum | VMAT Rectum | Rectum | VMAT | E2 | E2
2022-05-05 | 11 | Pelvis | Anus | Anus | VMAT | E1 | ELEKTA | 99.6 | 99.2 | 99.3 | After
2022-05-10 | 11 | Pelvis | Anus | Anus | VMAT | E2 | ELEKTA | 96.7 | 83.7 | 78.3 | After
| 11 | Pelvis | Anus | Anus | VMAT | E1 | E1
2022-05-10 | 11 | Pelvis | Anus | Anus | VMAT | E2 | E2 | 98.4 | 85.7 | 82.5 | After
2022-05-05 | 11 | Pelvis | Anus Boost | Anus | VMAT | E1 | ELEKTA | 99.7 | 99.7 | 99.5 | After
2022-05-10 | 11 | Pelvis | Anus Boost | Anus | VMAT | E2 | ELEKTA | 81.5 | 86.2 | 77.6 | After
| 11 | Pelvis | Anus Boost | Anus | VMAT | E1 | E1
2022-05-10 | 11 | Pelvis | Anus Boost | Anus | VMAT | E2 | E2 | 87.9 | 92.9 | 86.1 | After
2022-05-02 | 12 | Pelvis | Cervix Boost | Cervix | VMAT | E1 | ELEKTA | 98.8 | 98.9 | 98.2 | After
2022-05-10 | 12 | Pelvis | Cervix Boost | Cervix | VMAT | E2 | ELEKTA | 92.7 | 97.2 | 90.8 | After
| 12 | Pelvis | Cervix Boost | Cervix | VMAT | E1 | E1
| 12 | Pelvis | Cervix Boost | Cervix | VMAT | E2 | E2
2022-04-22 | 13 | Thorax | Esophagus | Esophagus | VMAT | E1 | ELEKTA | 98.9 | 97.3 | 94.2 | Before
2022-05-10 | 13 | Thorax | Esophagus | Esophagus | VMAT | E2 | ELEKTA | 95.3 | 85.3 | 80.4 | After
| 13 | Thorax | Esophagus | Esophagus | VMAT | E1 | E1
| 13 | Thorax | Esophagus | Esophagus | VMAT | E2 | E2
2022-04-12 | 14 | Thorax | L Lung | Lung | VMAT | E1 | ELEKTA | 97.1 | 98.1 | 99.8 | Before
| 14 | Thorax | L Lung | Lung | VMAT | E2 | ELEKTA
| 14 | Thorax | L Lung | Lung | VMAT | E1 | E1
| 14 | Thorax | L Lung | Lung | VMAT | E2 | E2
| 15 | Abdomen | Pancreas | Pancreas | VMAT | E1 | ELEKTA
2022-05-03 | 15 | Abdomen | Pancreas | Pancreas | VMAT | E2 | ELEKTA | 99.5 | 97.7 | 97.4 | After
| 15 | Abdomen | Pancreas | Pancreas | VMAT | E1 | E1
| 15 | Abdomen | Pancreas | Pancreas | VMAT | E2 | E2
2022-04-13 | 16 | Thorax2 | L Lung | Lung | VMAT | E1 | ELEKTA | 94.5 | 96.7 | 93.5 | Before
| 16 | Thorax2 | L Lung | Lung | VMAT | E2 | ELEKTA
| 16 | Thorax2 | L Lung | Lung | VMAT | E1 | E1
2022-05-10 | 16 | Thorax2 | L Lung | Lung | VMAT | E2 | E2 | 99.6 | 99.1 | 99.8 | After
2022-05-02 | 17 | Thorax | R Lung | Lung | SBRT | E1 | ELEKTA | 99.7 | 100 | 99.7 | After
2022-05-02 | 17 | Thorax | R Lung | Lung | SBRT | E2 | ELEKTA | 100 | 100 | 100 | After
| 17 | Thorax | R Lung | Lung | SBRT | E1 | E1
| 17 | Thorax | R Lung | Lung | SBRT | E2 | E2
2022-04-25 | 18 | Thorax1 | R Lung | Lung | VMAT | E1 | ELEKTA | 99.6 | 99.6 | 99.8 | Before
2022-04-27 | 18 | Thorax1 | R Lung | Lung | VMAT | E2 | ELEKTA | 99.2 | 98.5 | 99.2 | Before
| 18 | Thorax1 | R Lung | Lung | VMAT | E1 | E1
| 18 | Thorax1 | R Lung | Lung | VMAT | E2 | E2
2022-04-22 | 19 | Thorax | R Lung | Lung | VMAT | E1 | ELEKTA | 100 | 99.3 | 99.7 | Before
| 19 | Thorax | R Lung | Lung | VMAT | E2 | ELEKTA
| 19 | Thorax | R Lung | Lung | VMAT | E1 | E1
| 19 | Thorax | R Lung | Lung | VMAT | E2 | E2

## Analysis
Avg GPR
_ | Beam 1 | Beam 2 | Plan_
E1 | 98.755 | 99.18 | 98.36
ELEKTA | 98.755 | 99.18 | 98.36
Before | 98.4286 | 98.6 | 97.9429
After | 98.9308 | 99.4923 | 98.5846
E2 | 96.2211 | 94.7632 | 91.7895
E2 | 96.48 | 94.36 | 91.54
After | 96.48 | 94.36 | 91.54
ELEKTA | 96.1286 | 94.9071 | 91.8786
Before | 99.5 | 98.8 | 99.5
After | 95.5667 | 94.2583 | 90.6083
Overall Average | 97.5205 | 97.0282 | 95.159
This shape represents a table slicer. Table slicers are supported in Excel or later. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.
This shape represents a table slicer. Table slicers are supported in Excel or later. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.
This shape represents a table slicer. Table slicers are supported in Excel or later. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.
This shape represents a table slicer. Table slicers are supported in Excel or later. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.
This shape represents a slicer. Slicers are supported in Excel 2010 or later. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2003 or earlier, the slicer cannot be used.
This shape represents a slicer. Slicers are supported in Excel 2010 or later. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2003 or earlier, the slicer cannot be used.
This shape represents a slicer. Slicers are supported in Excel 2010 or later. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2003 or earlier, the slicer cannot be used.
This shape represents a slicer. Slicers are supported in Excel 2010 or later. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2003 or earlier, the slicer cannot be used.
This shape represents a slicer. Slicers are supported in Excel 2010 or later. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2003 or earlier, the slicer cannot be used.
This shape represents a slicer. Slicers are supported in Excel 2010 or later. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2003 or earlier, the slicer cannot be used.
This shape represents a slicer. Slicers are supported in Excel 2010 or later. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2003 or earlier, the slicer cannot be used.
This shape represents a slicer. Slicers are supported in Excel 2010 or later. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2003 or earlier, the slicer cannot be used.